✨Shimmer✨
User Guide

---
Contents
Contents
These are hyperlinks! ctrl-click to jump
- Installation
- Configuration
- Pre-shimming
- Connecting the clients
- Dynamic shimming
- Manual control
- Modifying Matlab
- Troubleshooting
NOTE: ALL COMMANDS ARE IN ITALICS AND BOLD AND SHOULD BE ENTERED EXACTLY AS GIVEN :)

---
Installation
Installation from scratch
- Where to find the files,
- Where to put them on the computers,

---
Installation
In the shared area, there is a shimmer_fresh_install directory (folder), this contains a fresh, complete copy of shimmer. Please do not overwrite this,
- Copy this directory onto the Skope (into C:\) and MRShim computer (into C:\skope\Documents\), rename the directory on each computer to “shimmer”,

---
Configuring network
Configuring network_description.ini
- How to use ifconfig or ipconfig,
- Which numbers to use,
- Where to put the finished file, and what to name it,

---
Configuring the network
- If you don’t know what a port, socket or address is please read the explanation in the troubleshooting section,
- This step requires knowing which clients are on which computer. The default configuration is:
- Skope computer: server, matlab, console1,
- MRShim laptop: mrshim, console2

---
Network Diagram
Todo: make a big Laura style diagram here.

---
Configuring the network
- For each computer:
- Open a command prompt/terminal,
- Run ifconfig and find the longest output under ‘ethernet’. Note down the address next to IPv4.
- In network_description.ini, change the address of each client on that computer, to the address you just noted down.

---
Configuring the network
- The output of ifconfig is shown here:
- Rename network_description_7t.ini to network_description.ini (this is the file name shimmer looks for),
- Copy the file (via the shared area) into the shimmer directory on each computer. Both computers need exactly the same file, as it lets them work out which client has connected.
Your address may be different!

---
Pre-shimming
Pre-shimming
- To make this recipe, you will need:
- A high-res field map,
- Arche’s coil_tmp.npy,
- A .scan file,
- Correct names and location, (labelled screenshot of directory)

---
Pre-shimming
- Perform a high resolution B0 map,
- Copy the .par and .rec into the shared area.
- Rename them both to field_map.par/field_map.rec,
- Drag the .<check> into Arche and compute the required shims, then close Arche,
- Copy the coil_tmp.npy file from the C:\mrshim\ directory into the shared area,

---
Pre-shimming
- Open the Skope software and perform an offresonance and positions calibration,
- Ensure you save the calibration and use it going forward <screenshot>,
- Collect some data, any data, any number of dynamics, any TR, we just need the .scan file which contains the position data,
- Copy the .scan file into the shared area,

---
Pre-shimming
- Copy the previous files from the shared directory into the ./shimmer/data/ directory on the Skope computer,
- In matlab_client.m, change the ScanID variable (line 72) to match the ID of the .scan file (the integer at the start),
- You are now ready to shim! (I hope)

---
Connecting the clients
Connecting the clients
- Navigate to shimmer directories,
- Start the server,
- Start the console client,
- Connect the matlab client,
- Start the MRShim client,

---
Skope computer:
- Open two command prompts,
- Navigate to shimmer directory:
- cd C:\Users\skope\Documents\shimmer\shimmer

---
Skope computer
- !! Ensure network description is correctly configured before starting the server !!
- In one prompt, start the server with:
  python .\shimming_server.py
- Should see:
Fun fact! On the Skope computer you can start the clients by double clicking on their .py file!

---
Skope computer
- In another prompt, start a console client with:
  python .\console_client.py
- When prompted, enter 1 for the console number,
- Should see:
  and:
Fun fact! Once the server has started, you can connect the clients in any order.

---
Skope computer
- In matlab, navigate to the matlab_client.m file,
- In the first block, ensure the correct shimmer_directory is uncommented
- Run the first 6 sections
  up to:

---
Skope computer
- Should see:
- And also the controls greyed out on Skope software,
- (To disconnect from the Skope server and restore manual control, run the final block:)

---
MRShim computer
- Navigate in file explorer to shimmer directory:
- Right click below file list
  and click “Open in Terminal”
- (Can also open a PowerShell
  and navigate with cd)

---
MRShim computer
!! Ensure JUPITER_PLUGGED_IN is set to True in mrshim_client.py if you want it to actually shim !!
Run python .\mrshim_client.py
- Should see:
- In this terminal, and at the server:

---
MRShim computer
If it is connected to Jupiter, you should see the message: <screenshot>

---
MRShim Computer
- Optional: you can start a second console client (console number 2) on the MRShim laptop,
- Open another terminal window the same way as before,
- Enter python .\console_client.py
- When prompted, tell it 2.

---
Skope computer
- Before you start shimming, the server console should look like this:
- You can also check which clients are connected by entering list at a console and checking the output on the server.

---
Start dynamic shimming
Dynamic shimming
- Start the shimming with relay sinope “!start”,
- Running the matlab script,
- Stop the shimming, relay sinope “!stop”,

---
Any console
- Enter at a [shimmer] prompt:
  relay sinope “!start”
- This executes the “!start” command from Sinope, which enables shimming.
- You should see on the MRShim terminal:

---
Matlab Client
- Now the fun part! Time to start receiving Skope data and sending currents.
- Run the next 2 sections of matlab code:
  %% get scan def and edit parameters
  %% start scan
- The program is now in a loop, use the stop button at the stop of Matlab to stop acquiring Skope data and calculating currents,
- (To pause shimming, use one of the consoles instead to send: relay sinope “!stop” so the shim currents are set to 0)
Note: matlab will change Skope’s directory to .\shimmer\data\skope_temp\, and that’s where all the data will go.

---
Matlab Client
- In the command window you should see a status section that continuously updates, (NOTE: not currently working)
- In the MRShim client, you should see something like (but with more currents, and probably not stopped):

---
Manual control
Manual control
- Client vs server commands,
- !start, !stop, !debug, !shim
- relay
- halt, list, status

---
Client vs Server Commands
- This is the most complicated part of using shimmer.
- Everything in shimmer is done with commands,
- Some commands are executed by the client, and some by the server,
- But all commands are entered at a console client (at the [shimmer]: prompt),
- The way you make a non-console client execute a command is by relaying the client command to it from the console.
This is because the input() function of Python stops anything else from being done, so we can’t accept input on a client we need to do something else (like apply shims).

---
Client vs Server Commands
ALL client commands start with an exclamation mark(!), so you could say they start with a bang. Any input not starting with an exclamation mark is sent to the server.
I want to:	What you should type at a [shimmer] prompt:
Execute a server command	Type the server command (no !) and hit enter,
Execute a client command as this client	Type the command and hit enter, remember client commands start with a bang(!).
Execute a client command as another client	Relay the command to that client. Type: relay <client name> <command> and hit enter. Remember the ! on the command!

---
Client commands
- The client commands. Start with a bang!
- Not every client can execute every command, e.g., consoles can’t shim. The name in brackets is the name of the client that can execute that command. <> denotes parameters.
Command (client)	What it does
!start (mrshim)	Enable shimming (send shims to Jupiter)
!stop (mrshim)	Disable shimming (sets shims to 0)
!shim <currents> (mrshim)	Repeats the <currents> until it fills 24 channels, then sends them to Jupiter, currents should be integer numbers of milliamps, separated by spaces.
!reset (mrshim)	Reset the connection to the Jupiter amplifiers without stopping shimming.
!status (mrshim)	Toggle printing Jupiter status messages whenever shims are applied.

---
Client commands
Command (client)	What it does
!echo <message> (all)	Prints <message> to the console. Used for debugging.
!debug (all)	Enables printing debug messages.
!reader (all) or no command at all at a console.	A special command, cause the client in question to wait for a packet from the server, or reads one it just got. Because input() is blocking, we use this to make a console read from the server. Otherwise we only read after sending a command.
Console:
Waiting for input
It’s a command
Handle it
Read response
Other packet arrives
!reader
Meanwhile...
Read it
There is no way to stop waiting for input, so the other packet would never get read until we tell it to.

---
Relay
- Syntax: relay <client name> <command>
- The relay command causes the client named to execute the command, just as though it had typed it itself,
- The client names are those in network_description.ini and also those that appear at the console when they connect.
- It is a server command (server routes the packet to the client), so there’s no bang, but if <command> is a client command, it does need the bang, because it is dealt with just as though it had been typed in.
- Because input() is blocking, we can’t have an input() on the mrshim client (it needs to be shimming) so this is how we get commands to it: relay mrshim <command>

---
Server commands
The server commands. No ! to start.
Command (client)	What it does
list	Lists connect clients and their addresses,
status	Prints some server status info.
disconnect	Disconnects whichever client sent this.
halt	Disconnects all clients then stops the server. The cleanest way to close.
relay <client> <command>	Relays a command to another client, causing it to execute it itself.

---
Modifying the matlab
Modifying the Matlab
- To ask Skope to send different data types, change the port number on line 61 to the corresponding port for that data. See the table on line 45 (both in matlab_client.m)
- To change the data processing, modify the function in calculate_currents.m
Change this number to one of these

---
Troubleshooting Directory
There is an explanation of sockets and ports on the next slide.
- Restarted shimmer and now the clients won’t connect
- Clients still won’t connect! (computer restart may help)
- It isn’t connecting to Jupiter

---
Sockets, Ports and Addresses
- A socket is an abstraction representing the connection between two computers. They are created and managed by the operating system.
- An address belongs to each computer on a network. Like a building address. It looks like 128.128.64.64
- A single address can have many ports on it (each corresponding to a socket), like flats in the building. Port numbers are usually arbitrary, but those lower than 4000 are reserved, and 8000 is special. They look like: :12345
- So the full address (yes we use the same word) of a socket looks like 128.128.64.64:12345 and this is what the other computer connects to.
- Shimmer stores the address and ports of each client/computer in the network_configuration.ini file.

---
Restarted shimmer and now clients won’t connect
- Cause: operating system is keeping sockets in the TIME_WAIT state,
- Solution: wait for them to be freed (can take up to 2 minutes)
- Check by running netstat at a windows (not shimmer) command prompt and seeing the status of the :2500x ports. If they are TIME_WAIT then they won’t connect. If they don’t appear in the list then they are free. If they are ESTABLISHED then something is using them (probably shimmer).

---
Output of Netstat
Clients will connect:
Clients won’t connect:

---
Clients still won’t connect!
- Cause: the IPs may have changed or the ports may be in use by another process (somehow)
- Ports:
  - Fix: first, close shimmer and wait five minutes. Then run netstat in a command prompt and look at the output. If any of the :2500x ports are in use then they are being used by another process and shimmer won’t be able to use them.
  - Change the port numbers in the network_configuration.ini file to any other set of numbers. Ensure you copy this new file to all the machines.

---
Clients still won’t connect!
- Cause: the IPs may have changed or the ports may be in use by another process (somehow)
- IPs:
  - At each computer, run ifconfig at a windows command prompt and find the output that has the most lines under the word ethernet.
  - Find the line saying IPv4 and copy that number down,
  - Check it matches the address entry in network_description.ini for each client running on that computer.
  - See configuration section of this guide

---
It isn’t connecting to Jupiter
- Check Jupiter is plugged in,
- Check the JUPITER_PLUGGED_IN variable at the top of mrshim_client.py is set to True,
- Check the device name in sinope and ensure it matches that in mrshim_client.py, self.jupiter_ifname (line 46)
- Ask Mags